## Year_2000_Compliance_Testing
Last Modification Date: | 12th May, 1998
MANUFACTURER | MODEL | EQUIPMENT TYPE | SOFTWARE VERSION | Y2000 Power On | Y2000 Power Off | Y2001 Power On | Y2001 Power Off | 30/10/2000 Monday? | Leap Year OK? | 9/9/99 OK? | Manfacturer/Vendor notified? | Manufacturer/Vendor Aware? | Manufacturer Comments Available? | Cost | Tested by | Date of Testing | Comment
3M | AVI 470 | Pump, Infusion, Volumetric | 051058 Rev A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | No |  | Hunter Area Health Service | 04/11/1997 | No date functions.
3M | AVI 275 | Pump, Infusion, Volumetric | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Abbott Laboratories | Flexiflo Companion | Pump, Infusion, Enteral | Not reported | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Yes |  | Hunter Area Health Service | 12/11/1997 | Unit is Y2K compliant. See Abbott's web page at http://www.abbott.com/year2k/y2kprodcompl.htm
Abbott Laboratories | Lifecare 4200 | Pump, Infusion, PCA | V 5.21 EUK | Pass | Pass | Pass | Pass | N/A | N/A |  |  |  | Yes |  | Hunter Area Health Service | 21/07/1997 | Date clicks over OK but can only set date in range 87 to 99. See Abbott's web page at http://www.abbott.com/year2k/y2kprodcompl.htm
Abbott Laboratories | PMP | Pump, Infusion, PCA | ? | Pass | Pass | Pass | Pass | N/A | N/A |  |  |  | Yes |  | Hunter Area Health Service | 18/07/1997 | Unit is Y2K compliant. See Abbott's web page at http://www.abbott.com/year2k/y2kprodcompl.htm
Abbott Laboratories | Lifecare 4200 | Pump, Infusion, PCA | V 6.13 EUK | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | Hunter Area Health Service | 18/02/1998 | Unit is Y2K compliant. See Abbott's web page at http://www.abbott.com/year2k/y2kprodcompl.htm
Abbott Laboratories | 50596-27-01 (Flexiflo Quantum) | Pump, Infusion, Enteral | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 24/02/1998 | No date / time functions. Unit is Y2K compliant. See Abbott's web page at http://www.abbott.com/year2k/y2kprodcompl.htm
Abbott Laboratories | Omniflow 4000+ | Pump, Infusion | Control rev 1.16dom Pump rev 1.16 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | No | N/A | Yes | N/A | New England Medical Center - Boston | 29/04/1998 | All tests OK. Unit is Y2K compliant.
Aesculap | GA 704 | Saw, Plaster | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Air-Shields | Isolette C100/200-2E | Incubator, Infant |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Air-Shields | Isolette C2HS-1C | Incubator, Infant | Rev 1.341 1.06 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. The RS232 connector provides output of operating parameters only. There is no date exchange through this connection.
Air-Shields | PT533-1E | Light, Ultraviolet, Phototherapy | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Air-Shields | PT533-2E | Light, Ultraviolet, Phototherapy | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Air-Shields | PM78-1E | Warmer, Overhead, Infant |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Air-Shields | C86 | Incubator, Infant |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Allersearch | VitalAir | Pump, Air, Nebuliser | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Ameda | Egnell Compact 25 | Pump, Suction | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Ameda | Egnell Universal 45 | Pump, Suction | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
American Medical Systems | K-20-D | Hyperthermia Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
American Pharmaseal | EC300 | Hyperthermia Unit, Control Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | No date / time functions.
American Pharmaseal | RK325 | Heater / Cooler | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 13/03/1998 | No date functions.
Anax | TCL15 | Light Source / Cautery Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Art Inc. | Cardio Lab | Electrophysiology System | V3.05 | Fail | Fail | Fail | Fail | N/A | Pass | N/A | Yes | No | No |  | St Vincent's Hospital - Sydney | 11/02/1998 | PC Based System, Laser Printer Reports Supplier: Micromedical 07:55739777 ; 01-01-2000 reported as 01-01-100 ; 01-01-2001 reported as 01-01-101 ; 01-01-1999 reported as 01-01-99.
ATL | HDI3000 | Ultrasound, Cardiac | 4252-0791-08 039.08 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 08/04/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Atom Medical | TC-500 | Incubator, Infant | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Atom Medical | OX-60 | Monitor, Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Augustine Medical | Bair Hugger 500/OR | Hyperthermia Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Augustine Medical | Bair Hugger 505 | Hyperthermia Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Augustine Medical | Bair Hugger 250 | Hyperthermia Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Australian Medical Engineering | AME/THREE | Pump, Intermittent Pressure Therapy Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
B Braun | Stimuplex S | Stimulator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
B&K Medical | 2001 | Ultrasonic Scanner | VP2698 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | N/A |  | No |  | Hunter Area Health Service | 03/11/1997
Baxter | 6301 | Pump, Infusion, Volumetric | Rev C | Pass | Pass | Pass | Pass | N/A | N/A |  |  |  | No |  | Hunter Area Health Service | 28/07/1997 | OK to 2020
Baxter | 6201 | Pump, Infusion, Volumetric | Rev C | Pass | Pass | Pass | Pass | N/A | N/A |  |  |  | No |  | Hunter Area Health Service | 28/07/1997 | OK to 2020
Baxter | BW-5 (Fenwal) | Blood Warmer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/02/1998 | No date / time functions.
Baxter | 6201 | Pump, Infusion, Volumetric | 1.12 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 19/02/1998 | Maximum valid year is 2020.
Baxter | OXYSAT SM-0200 | Monitor, Oxygen Saturation | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Bayer | Clinitek 50 | Analyser, Urine | 01.03 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No | N/A | Hunter Area Health Service | 29/04/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
BCI International | 71000A2 Oxycheck | Monitor, Pulse Oximeter | Ver 15, Jan 9 1997 #1819 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | No | N/A | Hunter Area Health Service | 11/12/1997
Bear Medical Systems | Bear 5 | Ventilator | Release 10 | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No | N/A | No |  | Hunter Area Health Service | 16/03/1998 | 28-Feb-00 rollsover to 01-Mar-00. The unit does record the "day of the week", though it is only user configurable, i.e. the system will roll it over appropriately but any date can be set to any "day of the week".
Bear Medical Systems | Bear Cub | Ventilator, Infant |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Bear Medical Systems | CEM | Enhancement Module, Ventilator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Bear Medical Systems | NVM-1 | Monitor, Ventilator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Biomec | 527B | Stimulator, Calf Muscle |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Bio-Med Devices | MVP-10 | Ventilator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | Gas Unit - No date / time functions.
Biosearch | 3600 | Pump, Infusion, Enteral | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | St Vincent's Hospital - Sydney | 01/05/1998 | Unit is Y2K compliant.
Bird | 10067 | Blender, Air / Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - No date / time functions.
Bird | 15365A (Avian) | Ventilator, Transport | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 13/03/1998 | No date functions.
Blease | Datum | Vaporiser, Anaesthetic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | Gas Unit - No date / time functions.
Bourns | BP200 | Ventilator, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Branson | 8200 | Ultrasonic Cleaner |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Carl Zeiss | 305363 | Microscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Carl Zeiss | 308248 | Microscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Carl Zeiss | S21 | Microscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Catalyst Research | Miniox III | Monitor, Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
CIG | 520524 | Monitor, Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
CIG | 520526 | Monitor, Circuit Disconnect | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
CIG | Midget 3 | Anaesthetic Machine | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - no date / time functions.
Cincinatti Sub Zero | Hemotherm 400M | Heater / Cooler | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Clary | Onguard UPS2-1K-2G-SBS-SP | Uninterruptable Power Supply | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Clements | Series 15 | Suction Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Clements | HSU | Suction Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - No date / time functions.
Clements | Compact | Suction Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | No date / time functions.
Clements | SUC82660 | Pump, Suction | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Cobe | 043650-001 | Computerised Perfusion Controller |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | Unknown | No |  | Hunter Area Health Service | 14/10/1997 | Unit uses 2 digit year. Unit is Y2K compliant.
Cobe | 043660-000 | Monitor, Computerised Perfusion Control |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 03/04/1998 | Unit interfaces with Cobe 043650 (computerised controller - tested previously OK). Unit is Y2K compliant
Cobe | 043600-001 | Pump, Blood |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 03/04/1998 | No date functions. Unit is Y2K compliant.
Cobe | 043049-001 | Monitor, Temperature, Time |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 03/04/1998 | This unit maintains the time, day and month, but no year information is stored, hence the rollover tests are not applicable.
Cobe | 043630-001 | Air Emboli Protection System | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 30/03/1998 | No date functions. Unit is Y2K compliant.
Codman | 82-6605 | Control Unit, Monitor, Neuro | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Cook | LINS-1000 | Insufflator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date function. The data connections on the back panel loop data from I/P to O/P, plus some outputs from microprocessor board. OK.
Cooper Lasersonics | CUSA NS-100 | Aspirator, Ultrasonic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Corometrics | 220 | pH Meter | Unknown | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No | N/A | No |  | Illawarra Area Health Service | 06/12/1997 | Every month is 31 days - need to set the 1st day each month - clocks over to 00 and 01
Corometrics | 116 | Monitor, Ultrasound, Fetal | 3.10 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 17/02/1998 | 31-12-99 rolls over to 1-1-0
Corometrics | 322 | Transmitter, Telemetry, Fetal | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/02/1998 | No date / time functions.
Corometrics | 321 | Receiver, Telemetry, Fetal | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/02/1998 | No date / time functions.
Corometrics | 116 | Monitor, Ultrasound, Fetal | 3.31 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 18/02/1998 | 31-12-99 rolls over to 1-1-0
Criticare Systems, Inc. | 602-13 | Monitor, ETCO2, SpO2 | M1.0, S1.0 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No | N/A | Prince of Wales Hospital - Sydney | 02/12/1997
Criticare Systems, Inc. | 507O | Monitor, NIBP, SpO2 | Rev M1.4 1995 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No |  | No |  | Hunter Area Health Service | 14/10/1997
Criticare Systems, Inc. | 507O | Monitor, NIBP, SpO2 | M1.3 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 12/03/1998 | This unit uses a 2 digit year, i.e. 99 rolls over to 00.
Criticare Systems, Inc. | 503 | Monitor, Pulse Oximeter |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 12/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Cuda | M2-150-T | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Cuda | M2-150 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Datex | Satlite | Monitor, Pulse Oximeter | 875314-1 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | No | N/A | Illawarra Area Health Service | 04/12/1997 | 24 hr clock only - no date function
Datex | Satlite Trans | Monitor, Pulse Oximeter | PR 877014 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | No | N/A | Illawarra Area Health Service | 04/12/1997 | No date functions.
Datex | Cardiocap CH-2S | Anaesthesia Monitor | PR 882593-3.2 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Illawarra Area Health Service | 09/12/1997
Datex | Cardiocap CM-104-21-01 | Monitor, ECG, Press, NIBP | 877009-1/21.11.89 | Pass | Pass | Pass | Pass | N/A | Pass | Not tested | No | N/A | No |  | Hunter Area Health Service | 05/03/1998 | Can not determine correct date rollover while the unit is on. Need to access Engineering menu (via a power off - power on procedure) to determine the date.
Davol | System 5000 | Electrosurgical Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Davol | System 2000BP | Electrosurgical Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Desoutter | CC1 | Saw, Plaster | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Devilbiss | Ultra-Neb 2000 | Nebuliser, Ultrasonic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions.
Downs | MK II | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Downs | MKI | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Drager | PM8050cd | Monitor, Patient Airway |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | The Queen Elizabeth Hospital - Adelaide | 29/10/1997
Drager | Oxylog | Ventilator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - No date / time functions.
Drager | Evita 4 | Ventilator | 02.20 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | No | N/A | N/A | N/A | St Vincent's Hospital - Sydney | 12/05/1998 | Unit uses 2 digit year. Unit is Y2K compliant (windowed compliance).
Dyonics | Autobrite II | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 08/10/1997 | No programmable logic or real time clock.
Dyonics | 750 PAL | Video Camera Unit | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions. The manuals do not refer to any date based information, with regard to possible date exchange through any of the rear panel connectors.
Dyonics | Digital PAL Video Camera | Video Camera Unit | Ver A.00 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Dyonics | DyoPneumatic 15 | Insufflator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Electa/Selector | 1522000 | Aspirator, Ultrasonic | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | No | N/A | Hunter Area Health Service | 10/11/1997 | Dumb device - no clock or calendar
EME | TC 2-64 | Monitor, Ultrasound, Doppler, Transcranial | TC2.5B-E | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | No date / time functions. The D-connectors on the rear of the unit do not facilitate the exchange any date info.
Fisher and Paykel | MR620 | Humidifier | A20 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date function. Serial data connection on base of unit is an output only.
Fisher and Paykel | MR512 | Humidifier | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Fisher and Paykel | MR600 | Humidifier | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions.
Fisher and Paykel | MR500 | Humidifier |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Fisher and Paykel | MR730 | Humidifier |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Fisher and Paykel | MR700 | Humidifier |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Fisher and Paykel | MR428 | Humidifier |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Fisher and Paykel | MR720 | Humidifier |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/04/1998 | No date functions. Unit is Y2K compliant.
Fresenius | 2008A | Dialysis Machine | V3.2 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | Yes | N/A | No | N/A | South Eastern Sydney Area Health Service | 01/04/1998 | Unit is Y2K compliant.
Fresenius | 4008E | Dialysis Machine | V3.2 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | N/A | N/A | No | N/A | South Eastern Sydney Area Health Service | 01/04/1998 | Unit is Y2K compliant.
Gambro Healthcare | Prisma CFM 018080000 | Dialysis Machine | R01_02_A | Pass | Pass | Pass | Pass | N/A | Pass | Pass | N/A | N/A | Yes |  | Hunter Area Health Service | 12/02/1998
Gambro Healthcare | Prismatherm FW240 | Warmer, Dialysate | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 12/02/1998 | No date / time functions.
Gambro Healthcare | Prisma | Dialysis Machine | R02_04_A3 | Pass | Pass | Pass | Pass | N/A | Pass |  |  |  | Yes |  | St Vincent's Hospital - Sydney | 27/02/1998 | There is a real time clock that logs time between services. It then gives an audible alert, which can be overridden by the operator until it is convenient to perform the service.
Gambro Healthcare | AK-200 | Dialysis Machine | V1.11 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | Yes | N/A | Yes | N/A | South Eastern Sydney Area Health Service | 01/04/1998 | Unit is Y2K compliant.
Gambro Healthcare | AK-100 | Dialysis Machine | V3.71 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | South Eastern Sydney Area Health Service | 01/04/1998 | No chronological date involved. Unit is Y2K compliant.
Gaymar | Thermocare | Hyperthermia Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Gorman-Rupp | K-12 | Hyperthermia Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Graseby | PCAS | Pump, Infusion, PCA | Not reported | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Yes |  | Hunter Area Health Service | 12/11/1997
Graseby | 3300 | Pump, Infusion, PCA | 4.20E | Pass | Pass | Pass | Pass | N/A | Pass | Pass | N/A | N/A | Yes |  | St Vincent's Hospital - Sydney | 21/01/1998
Graseby | MS 16A | Pump, Infusion, PCA | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Graseby | 3300 | Pump, Infusion, PCA | 4.30E | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | Hunter Area Health Service | 23/02/1998
Graseby | MR10 | Monitor, Respiration, Apnoea |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Graseby | MS26 | Pump, Infusion, Syringe |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Hawksley | 6/40 | Blanket Warmer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 03/11/1997 | No programmable logic or real time clocks.
Hawksley | MK4 | Blanket Warmer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 03/11/1997 | No programmable logic or real time clocks.
Hawksley | MK5 | Blanket Warmer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 03/11/1997 | No programmable logic or real time clocks.
Hawksley | MK7 | Blanket Warmer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 03/11/1997 | No programmable logic or real time clocks.
Healthdyne | 10100 | Monitor, Pressure | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date / time functions. Unit is Y2K compliant
Hemocue | Hemocue | Monitor, Haemoglobin | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 16/03/1998 | No date functions.
Hewlett Packard | M1702A | Recorder, Electrocardiograph | Not specified | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | The Queen Elizabeth Hospital - Adelaide | 28/10/1997
Hewlett Packard | 78560 | Monitor, Central | 82200E V7.0 | Pass | Pass | [7 empty] | Yes |  | St Vincent's Hospital - Sydney | ??/??/1997
Hewlett Packard | M1125A | Recorder, Bedside |  | Pass | [8 empty] | Yes |  | St Vincent's Hospital - Sydney | ??/??/1997
Hewlett Packard | Sonos 500 | Ultrasound, Cardiac | Rev K | Pass | [8 empty] | Yes |  | St Vincent's Hospital - Sydney | ??/??/1997
Hewlett Packard | 4745 (Pagewriter 2) | Electrocardiograph |  | Pass | Pass | Pass | Pass | N/A |  |  | No | N/A | Yes |  | North West Adelaide Health Service | 30/10/1997
Hewlett Packard | Series 50 XM | Monitor, Foetal/Maternal | B.01.06 M1350-6801U | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | Illawarra Area Health Service | 14/11/1997
Hewlett Packard | M1723A (Codemaster XL ) | Defibrillator/Monitor |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | Illawarra Area Health Service | 08/12/1997
Hewlett Packard | M1771A Pagewriter | Recorder, Electrocardiograph | A.05.06 English | Pass | Pass | Pass | Pass | Pass | Pass | Pass | No | N/A | Yes | N/A | Illawarra Area Health Service | 08/12/1997
Hewlett Packard | 78560A | Monitor, Central | Rev 2.0 | Pass | Pass | Fail | Fail | N/A | Fail | N/A | No | N/A | Yes |  | St Vincent's Hospital - Sydney | 27/01/1998 | Passes year 2000 rollover, but does not recognise 2000 as a leap year. Recordings have 2 digit year.
Hewlett Packard | Sonos 5500 | Ultrasound, Cardiac | A:0 | Pass | Pass | Pass | Pass | Pass | Pass | N/A | N/A | N/A | Yes |  | St Vincent's Hospital - Sydney | 10/11/1997
Hewlett Packard | 8040A | Monitor, Ultrasound, CTG | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 17/02/1998 | Unit only maintains clock. No date information utilised.
Hewlett Packard | 43120A | Defibrillator/Monitor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 20/03/1998 | No date / time functions on the unit tested. Note that some of the defibrillators in the 43100A series do have date functions.
Hewlett Packard | 78354A | Monitor, Multiparameter |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Hewlett Packard | 77020A | Ultrasound, Cardiac | See Comments | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No | N/A | No | Unknown | Hunter Area Health Service | 07/04/1998 | Can NOT alter date and time settings on the 29 Feb 00 (produces Error Code 00254). Note that functiionality is NOT affected (Carried out POWER-ON and BASIC tests OK). S/W ver: CT/PRO O.0 CT/MEM O.0 BE O.0.2 SAIL O.0 BE/RST N.3.1 FE O.0.1 FE/RST O.0.1
Hewlett Packard | 77030A | Ultrasound, Cardiac | See Comments | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No | Yes | No | Unknown | Hunter Area Health Service | 23/04/1998 | Can NOT alter date and time settings on the 29 Feb 00 (produces Error 00254). Note that functiionality is NOT affected (Carried out POWER-ON and BASIC tests OK). S/W ver: CT/PRO N.3 CT/MEM N.3.1 BE N.3.2 SAIL .2 BE/RST N.3.1 FE N.3.1 FE/RST N.3.1
Hewlett Packard | 78352A | Monitor, Multiparameter |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 04/05/1998 | No date functions. Unit is Y2K compliant.
Hokanson | EC4 | Plethysmograph | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/03/1998 | No date / time functions.
Hokanson | E20 | Rapid Cuff Inflator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/03/1998 | No date / time functions.
Howorth Engineering | PCM-40 | Blanket Warmer |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 04/11/1997 | No programmable logic or real time clocks.
Hudson | 5577 | Monitor, Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Huntleigh | Audio Dopplex | Doppler Vascular Flowmeter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 11/11/1997 | No real time clocks
Huntleigh | DFS1000 | Sequential Compression Device | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 11/11/1997 | No programmable logic or real time clocks.
Huntleigh | Mini Dopplex | Doppler Vascular Flowmeter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 11/11/1997 | No real time clocks
Huntleigh | AC500/DVT | Sequential Compression Device | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 11/11/1997 | No programmable logic or real time clocks.
Huntleigh | NB 2000 (Nimbus II) | Mattress, Flotation | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 16/03/1998 | No date / time functions.
Huntleigh | Flowtron AC/200 | Pump, Intermittent Pressure Therapy Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Huntleigh | D920A | Monitor, Ultrasound, Fetal |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 29/04/1998 | No date / time functions. Unit is Y2K compliant
Imed | Gemini PC-1 | Pump, Infusion, Volumetric | 7.01 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | No |  | Hunter Area Health Service | 04/11/1997
Imed | PC-4 | Pump, Infusion | XL5.11.17.0 | Pass | Pass | Fail | Fail | Fail | Fail | Pass | Yes | N/A | No | Unknown | St Vincent's Hospital - Sydney | 01/05/1998 | Can not enter the year 00. (invalid year entered). Upgrade software to be available Sept / Oct 1998.
Infrasonics | Infant Star HFV | Ventilator, Infant |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Infrasonics | 606100 (Star Sync) | Synchroniser, Ventilator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date / time functions. The connectors on the rear of the unit do not facilitate the exchange any date info. The RS232 connector is an output for parameter values only. Unit is Y2K compliant.
Infrasonics | 1606004 (Star Sync) | Synchroniser, Ventilator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
International Technidyne Corporation | Hemochron 401 | Monitor, Blood Clotting Time | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 16/03/1998 | No date functions
Invivo | 4500 | Monitor, Pulse Oximeter |  | Pass | Pass | Pass | Pass | Pass | Pass | Not tested | N/A |  | No |  | North West Adelaide Health Service | 07/11/1997
Invivo | 4500 | Monitor, Pulse Oximeter | Rev OX06 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 20/02/1998
Irex | System II | Ultrasound, Cardiac | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 13/03/1998 | No date functions.
Ivac | Ivac 711 | Pump, Infusion, Syringe | 71X15007.04 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | No |  | Hunter Area Health Service | 04/11/1997 | No date functions.
Ivac | 591 | Pump, Infusion, Volumetric |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Johnson & Johnson Medical | DINAMAP Compact TS | Monitor, NIBP, SpO2, Temp | 8613-V4.1 8621-V3.1 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 02/03/1998 | Uses 4 digit year. Can set year anywhere from 1996 to 2099
Johnson & Johnson Medical | Dinamap 8100 | Monitor, NIBP | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date function. Communications port does not exchange date information (time only).
Johnson & Johnson Medical | 1846 SX | Monitor, NIBP |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Jostra | PRM 15-450 | Monitor, Pressure | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
JVC | TM-210-PS | Monitor, Video | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions.
Kay | 9100 | RhynoLaryngeal Stroboscope |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | The Queen Elizabeth Hospital - Adelaide | 22/01/1998
Keeler | ACU-14 | Cryosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Kendall | SCD 5325 | Sequential Compression Device | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date / time functions.
Kontron | KAAT-II | Balloon Pump, Intra Aortic | 3.02.3 | Pass | Pass | Pass | Pass | N/A | Pass |  | N/A |  | No |  | Hunter Area Health Service | 20/08/1997
Kontron | Microgas 7640 Mk2 | Monitor, TCPO2 | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Kontron | 7650-500 | Monitor, TCPO2 / TCCO2 | 2.01 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | N/A | N/A | No | N/A | Hunter Area Health Service | 27/04/1998 | Unit is Y2K compliant.
Laerdal | Heartstart 3000 | Defibrillator/Monitor | G | Pass | Pass | Fail | Fail | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 21/07/1997 | 2000 rolls over to 1901
Laerdal | LSU | Pump, Suction | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Laerdal | Heartstart 3000 | Defibrillator/Monitor | H | Pass |  | Pass | Pass | N/A | Fail | N/A | No | N/A | No |  | St Vincent's Hospital - Sydney | 10/03/1998 | Does not recognise 29 Feb 2000.
Life Support Products | Autovent 2000 | Ventilator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 13/03/1998 | Gas Unit - no date / time functions.
Lifecare | PLV-100 | Ventilator | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Linvatec | 8425A | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Loosco | Mark 2 | Ventilator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Luxtec | 3150AS | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 30/03/1998 | No date functions. Unit is Y2K compliant.
Mallincrodt | Warmtouch 5700 | Blanket Warmer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 07/10/1997 | No programmable logic or real time clock.
Marquette Medical Systems | MAX-1 | Stress ECG testing system | 001F | Pass | Pass | Pass | Pass | Pass | Pass | Pass | No | N/A | No | N/A | The Queen Elizabeth Hospital - Adelaide | 03/12/1997
Marquette Medical Systems | 7010 | Monitor, Mainframe | 83605/612-103 (1988) | Fail | [8 empty] | No |  | St Vincent's Hospital - Sydney | ??/??/1997 | System (MBX) time used on paper recordings. Fails 1 Jan 1900
Marquette Medical Systems | 7500 | Monitor, Central | MBX8325>64-002D | Pass | [8 empty] | No |  | St Vincent's Hospital - Sydney | ??/??/1997 | System controller software
Marquette Medical Systems | Solar 8000/Tram 450Sl/TramRac 4A | Monitor, Multiparameter | 4030 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No | N/A | Illawarra Area Health Service | 20/11/1997
Marquette Medical Systems | MAC-15 | Monitor, ECG | 006B | Fail | Fail | Fail | Fail | N/A | Fail |  | No |  | No |  | St Vincent's Hospital - Sydney | 26/03/1998 | With power on, the unit displays the correct rollover date, i.e. 01 Jan 00, though on a print out, the date is displayed as 03 Mar A6. After power off and on again, the system date reverts to 01 Jan 80.
Marquette Medical Systems | MAC-LAB | Monitoring System, Catheter Lab | Ver 5D | Fail |  | Fail |  |  | Pass |  | Yes |  | No | $0, Upgrade to S/W version 6F. | St Vincent's Hospital - Sydney | 25/02/1998 | Rolls over from 23:59 31 Dec 99 to 00:00 1 Jan 10. The year "10" does not advance at 2001. Also the year cannot be manually changed to "00".
Marquette Medical Systems | CASE 15 | Stress ECG testing system | EPROM 004B CASE 006C | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 18/03/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Marquette Medical Systems | Eagle 3100 | Monitor |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Prince of Wales Hospital - Sydney | 23/01/1998 | Unit is Y2K compliant.
Marquette Medical Systems | Solar 7000 | Monitor, Modular | Tram407027-02204Apr95 7F3A Main Proc 414482-001 1A | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Prince of Wales Hospital - Sydney | 23/01/1998 | Unit is Y2K compliant.
Marquette Medical Systems | Eagle 4000 | Monitor | 409263-0175B 409259-0015B 409258-0025B | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Prince of Wales Hospital - Sydney | 23/01/1998 | Unit is Y2K compliant.
Marquette Medical Systems | 1900 | Treadmill |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 25/03/1998 | This unit interfaces with Marquette controlling devices. eg. CASE-15, MAX-1, etc. There is no date information exchanged through either the serial or parallel data ports.
Marquette Medical Systems | Eagle 3000 | Monitor, Multiparameter | 415230-0043A 414957-0043A 414959-0043A 415257-0011 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 07/04/1998 | Unit uses 4 digit year. Unit is Y2K compliant.
Marquette Medical Systems | 1950 | Monitor, NIBP |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 23/04/1998 | Unit uses 4 digit year. Unit is Y2K compliant.
Marquette Medical Systems | Seer / XT | Monitor, ECG, Holter |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 27/04/1998 | Unit has a 24 hour time function, but no date functions. Unit is Y2K compliant.
MDE | Escort II+ | Monitor, ECG, Resp, NIBP, SpO2, Temp | Main 1.10.00 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | Hunter Area Health Service | 04/11/1997 | Unit is Y2K compliant.
MDE | 20401 Prism EL | Monitor, Multiparameter | 02.02.00 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 19/03/1998 | Unit is Y2K compliant.
Mead Johnson | Maxi-myst | Pump, Air, Nebuliser | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 29/04/1998 | No date / time functions. Unit is Y2K compliant
Medasonics | FP3B | Monitor, Ultrasound, Fetal | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Medcor | IW802 | Warmer, Overhead, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 16/03/1998 | No date / time functions.
Medela | 0162003 | Pump, Breast | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date / time functions. Unit is Y2K compliant
Medelec | DG Examiner | Electroencephalograph | V5.07 | Pass | Pass | Not yet tested | Not yet tested | N/A | Not yet tested | Not yet tested |  |  | No |  | St Vincent's Hospital - Sydney | 20/08/1997 | Uses PC system date, dates on recordings are 2 digit
Medelec | PREMIER Plus | Electromyograph | VD.00 | Pass | Fail | Pass | Pass | Pass | Pass | Not tested |  |  | No |  | St Vincent's Hospital - Sydney | 20/08/1997 | Fails power off roll-over, PC based, PC has 4 digit date but EMG software has 2 digit date - goes to 1 digit for > 2000. (Fails - Goes to 04/01/1980)
Medequip | MQ775 | Warmer, Overhead, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Medgraphics | 1070 Series 2 | Pulmonary Function Analyser | Not known | Pass | Not tested | Pass | Not tested | Not tested | Not tested | Not tested | No |  | No |  | Hunter Area Health Service | 18/09/1996 | Date change Ok but because of application software patient results for > 2000 are listed before those <2000 in test table
Medison Pty Ltd | Sonoace 600 | Ultrasonic Scanner | SA 600 V1.40 | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No | Unknown | No |  | Hunter Area Health Service | 01/12/1997 | Leap year OK with power off but not power on - solution leave off overnight from 28/2 to 29/2/2000
Medrad | Mark V plus | Injector, Dye | MV7.0 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date / time functions.
Medrad | Mark IV | Injector, Dye |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date functions.
Medtel | HS5 | Monitor, Multiparameter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 27/04/1998 | No date functions. Unit is Y2K compliant.
Medtel | HS18 | Monitor, Multiparameter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 27/04/1998 | No date functions. Unit is Y2K compliant.
Medtel | HS40 | Monitor, Multiparameter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/04/1998 | No date functions. Unit is Y2K compliant.
Medtel | Commander | Monitor, Multiparameter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 27/04/1998 | No date functions. Unit is Y2K compliant.
Medtronic | 5330 | Pacemaker, External | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Medtronic | 5346 | Pacemaker, External | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 16/03/1998 | No date functions.
Medtronic | 5375 | Pacemaker, External | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/03/1998 | No date / time functions.
Medtronic | 5345 | Pacemaker, External | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/03/1998 | No date functions.
Medtronic | CardioCare 4220 | Monitor, ECG, Transtelephonic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/03/1998 | No date functions.
Medtronic | 0610 CardioRhythm | Ablation System |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date functions. The DB9 connector on the rear facilitates external monitoing of system parameters. There is no date exchange through this connection.
Medtronic | 5328 | Stimulator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date functions.
Mira | TR4000 | Diathermy Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Mira | OS3000 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Mortara | ELI200 | Recorder, Electrocardiograph | 7.93 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | ??/12/1997 | Tested at JHH and Wollongong
Mortara | ELI100 | Recorder, Electrocardiograph | Unknown | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 06/03/1998 | Uses a 2 digit year. 99 rolls over to 00.
Mortara | 30300-001-2001 (Portrait) | Recorder, Electrocardiograph | 9.97 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 12/03/1998 | This unit uses a 2 digit year, i.e. 99 rolls over to 00.
Mutech Instruments | 271E (Kabtronics) | Stimulator, ECT | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Nascor | Series 2A | Monitor, Transport |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 06/04/1998 | No date functions. Unit is Y2K compliant.
Nellcor | N-100 | Monitor, Pulse Oximeter | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Nellcor | Ultra Cap | Monitor, Pulse Oximeter | 1.08 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | N/A | N/A | No | N/A | St Vincent's Hospital - Sydney | 09/04/1998 | There is a printer port on this device which was not tested. As a stand alone unit the device is Y2K compliant.
Nellcor Puritan Bennett | 7200ae | Ventilator | TEM 26300-85-V B5Z | Pass | Pass | Pass | Pass | N/A | Pass | Pass | NO |  | No |  | Hunter Area Health Service | 13/10/1997
Nellcor Puritan Bennett | 7200a | Ventilator | 26160-85-R ARR | Fail | Fail | Fail | Fail | N/A | Not tested | Pass | No |  | No |  | Hunter Area Health Service | 13/10/1997
Nellcor Puritan Bennett | NPB-195 | Monitor, Pulse Oximeter | 2.0.1.0 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 24/03/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Nellcor Puritan Bennett | 7200 | Ventilator | SPI 26160-85-R AAB | Fail | Fail | Fail | Fail | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 16/03/1998 | The year field fails to roll over past 99, but the year can be manually set to the correct year (i.e. 00, 01, etc..).
Nellcor Puritan Bennett | NPB-195 | Monitor, Pulse Oximeter | 2.0.3.0 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | N/A | N/A | Yes | N/A | St Vincent's Hospital - Sydney | 02/04/1998 | Unit is Y2K compliant.
Nellcor Puritan Bennett | 1400005 Star Track | Monitor, Ventilator | 9600 900 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 07/04/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Nellcor Puritan Bennett | Infant Star 950 | Ventilator, Infant |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. The connectors on the rear of the unit do not facilitate the exchange any date info. The RS232 connector is an output for parameter values only. Unit is Y2K compliant.
Nellcor Puritan Bennett | NPB-295 | Monitor, Pulse Oximeter | 2.0.9.0 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 29/04/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Nihon Kohden | TEC-7200 A/B | Defibrillator/Monitor/Recorder |  | Pass | Pass | N/A | N/A | N/A | N/A | N/A | No | N/A | No |  | St Vincent's Hospital - Sydney | 14/10/1997 | Nihon Kohden claims Y2K compliance for this device.
Nihon Kohden | VC-680G | Monitor, Multitrace | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 03/04/1998 | No date functions. Unit is Y2K compliant.
Nihon Kohden | TEC-7200 K | Defibrillator/Monitor |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 23/03/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Nihon Kohden | RMC-1100M | Monitor, Multitrace | CPU V2.2 GDC V1.1 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Nisco | J10 | Video Camera Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Nonin | 8600 | Monitor, Pulse Oximeter |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date functions. The D-connector on rear panel is recorder output only. There is no date exchange through this connection.
Novametrix | Oxypleth | Monitor, Pulse Oximeter | ENG 520a-25 May-17-93 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Illawarra Area Health Service | 27/11/1997
Novametrix | 515 | Monitor, Pulse Oximeter | 1.6 Jul 20, 90 | Fail | Fail | Fail | Fail | N/A | Pass | Pass | Yes | No | No |  | Illawarra Area Health Service | 05/12/1997 | Will not roll over for any year - need to manually set the date at the beginning of each year - othervise date comes up as "MO" - Note: This is only a problem if the unit is attached to a printer
Ohio | Air-Vac | Incubator, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Ohmeda | 3800 | Monitor, Pulse Oximeter | 3.000/03.000 | N/A | N/A | N/A | N/A | N/A | N/A |  |  |  | Yes |  | Hunter Area Health Service | 31/07/1997 | No clock function
Ohmeda | 2300 | Monitor, NIBP | 4.0 | Pass | Pass | Pass | Pass | N/A | Pass |  | N/A |  | Yes |  | Hunter Area Health Service | 12/09/1997
Ohmeda | Tec 3 | Vaporiser, Anaesthetic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 23/02/1998 | Gas Unit - No date / time functions.
Ohmeda | Tec 4 | Vaporiser, Anaesthetic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 23/02/1998 | Gas Unit - No date / time functions.
Ohmeda | Quantiflex | Anaesthetic Machine | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 04/03/1998 | Gas Unit - No date / time functions.
Ohmeda | Ohio 2001 | Warmer, Overhead, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Ohmeda | Midogas A90 | Blender, Oxygen, Nitrous Oxide | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/02/1998 | Gas Unit - No date / time functions.
Ohmeda | ISU | Suction Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - No date / time functions.
Ohmeda | 3700e | Monitor, Pulse Oximeter | 22 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 24/02/1998 | No date functions. No date information exchange through communications port.
Ohmeda | 3740 | Monitor, Pulse Oximeter | 10 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes |  | Hunter Area Health Service | 24/02/1998 | Date information from Digital Interface is output from unit only. (2 digit year ie. 2000 represented as 00)
Ohmeda | 7000 | Ventilator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Ohmeda | 4700 Oxicap | Monitor, Capnograph, Pulse Oximeter | Rev 7 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 27/03/1998 | Unit uses 4 digit year. Unit is Y2K compliant.
Olympus | CLK-4 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Olympus | LF-1 | Bronchoscope, Intubating | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Olympus | LF-2 | Bronchoscope, Intubating | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Olympus | CLK-3 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Olympus | BF Type P20 | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Organon Technika | Digistim III | Stimulator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Oxford | MR14 | Monitor, ECG, Holter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 06/03/1998 | No date functions.
Panasonic | AG-7350 | Video Cassette Recorder |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/04/1998 | No date functions. Unit is Y2K compliant.
Panasonic | AG-7330 | Video Cassette Recorder |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Penlon | PPV Sigma | Vaporiser, Anaesthetic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | Gas Unit - No date / time functions.
Pentax | EPM-3000 | Fibre Optic Light Source/Video processor | Not known | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No |  | No |  | Hunter Area Health Service | 06/11/1997
Pentax | LX-500A | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Pentax | LH-150 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Pentax | FNL-10S | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Pentax | FB-10X | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Pentax | FB-15X | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Pentax | FG-16X | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Pentax | EC-3400F | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Pentax | EG-2700 | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Pentax | EG-3800T | Endoscope | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Physio Control | CMS 8000 | Monitor, Central | 2.3 | Pass | Not tested | Pass | Not tested | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 24/11/1997
Physio Control | Lifepak 10 | Defibrillator/Monitor | 017J (System 0) | Pass | Pass | [7 empty] | No |  | St Vincent's Hospital - Sydney | ??/??/1997 | Year 2000 changeover is noted in the manual
Physio Control | Lifepak 9B | Defibrillator/Monitor | 1989, PCC | Pass | Pass | Pass | Pass | N/A | Pass | Not tested | No |  | No |  | Hunter Area Health Service | 14/10/1997
Physio Control | Lifepak 5 | Defibrillator/Monitor/Recorder | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 14/10/1997 | No programmable logic or real time clock.
Physio Control | Quik-Pace | Pacemaker, External | 3.10 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Physio Control | Lifepak 6S | Defibrillator/Monitor/Recorder | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Physio Control | Lifepak 10 | Defibrillator/Monitor | System 017 : Defib 23 : Chart/Pacer 15 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 24/03/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Physio Control | Lifepak 7 | Defibrillator/Monitor | pre 1985 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 08/05/1998 | The tested unit has no date functions. Unit is Y2K compliant (full compliance). Physio Control have noted that Lifepak 7 defibrillators manufactured since 1985 DO have date functions, and are NOT Y2K compliant.
Precision Electronic Products | SP CLUB | Ultrasound Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Protocol Systems, Inc. | Propaq 106 | Monitor, Transport | 5.02.00E | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No |  | Hunter Area Health Service | 04/12/1997
Protocol Systems, Inc. | Propaq Encore | Monitor, Transport | 1.30.00 | Pass | Pass | Pass | Pass | N/A | Fail | Pass | Yes | Yes | No |  | Illawarra Area Health Service | 05/12/1997 | Clock ticks over to 29/2/2000 but cannot enter time changes on that day - Function not impaired!
Protocol Systems, Inc. | 206EL | Monitor, Transport | 1.00.03 | Pass | Pass | Pass | Pass | N/A | Fail | Pass | Yes | Yes | No |  | Hunter Area Health Service | 13/03/1998 | Monitor rollsover to 29/02/00 OK, but can not set time on 29/02/00. Monitor stilll functions OK.
Protocol Systems, Inc. | Propaq 102 | Monitor, Transport | 4.09.00E | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | N/A | N/A | University Hospital Nijmegen - The Netherlands | 04/05/1998 | No day of week indication. Unit is Y2K compliant.
Protocol Systems, Inc. | Propaq Encore | Monitor, Multiparameter | 1.31.01 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | No | N/A | No | N/A | St Vincent's Hospital - Sydney | 28/04/1998 | Unit uses 2 digit year. Unit is Y2K compliant.
Radionics Inc | SCS-1 | Stereotaxy Computer | STAX-84K V4 | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No |  | No |  | Hunter Area Health Service | 25/09/1997 | Can reset time and date manually for leap year
Radionics Inc | SCS-1 | Stereotaxy Computer | UCLF V.1 | Pass | Pass | Pass | Pass | N/A | Fail | Pass | No |  | No |  | Hunter Area Health Service | 25/09/1997 | Can reset date and time manually for leap year
Radionics Inc | RFG-6 | RF Lesion Generator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
RB | CLS150 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Respironics | 302220 | Monitor, Circuit Disconnect | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 13/03/1998 | No date functions.
Seca | 727 | Scales, Electronic, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Sechrist | 3500HL | Blender, Air / Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - no date / time functions.
Sedco | 2510 | Tape Eraser | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Sensormedics | 2900 | Metabolic Cart | Unknown | Pass | Fail | Not tested | Not tested | Not tested | Not tested | Not tested | No | No | No |  | Hunter Area Health Service | 09/04/1998 | This unit is use an IMB PS/2 PC, running MS-DOS. The PC section is the only section of the device that utilises date-based information, and it fails Y2K compliance testing, though the primary function of the device as an analyser is not impaired.
Servox | Inton | Speech Device, Electronic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date / time functions.
Sherwood Medical | Genius 3000A | Thermometer, Electronic | Not Known | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | No |  | Hunter Area Health Service | 13/10/1997
Sherwood Medical | Kangaroo 330 | Pump, Infusion, Enteral | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 13/10/1997 | No programmable logic or real time clock.
Sherwood Medical | Kangaroo 22 | Pump, Infusion, Enteral | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 07/10/1997 | No computer technology.
Sherwood Medical | Kangaroo 224 | Pump, Infusion, Enteral | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 09/12/1997 | No programmable logic or real time clock.
Sherwood Medical | Kangaroo 324 | Pump, Infusion, Enteral | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 24/02/1998 | No date functions.
Siemens | 121C | Monitor, Central | Not stated | Fail | Not tested | Not tested | Not tested |  |  | Pass | Not reported |  | No |  | Repatriation General Hospital - South Australia | 20/11/1997 | Central monitor changes over to 2000 initially but reverts to 1900 after a few minutes - bedside monitors networked to the central also revert to 1990 but are OK when tested alone
Siemens | 1281C, 981C & 732 | Monitor, Bedside | Unknown | Pass | Pass | Pass | Pass | Pass | Pass | Pass | Unknown |  | No |  | Repatriation General Hospital - South Australia | 20/11/1997 | All bedside monitors above pass the above tests when operating stand-alone but revert to 1900 when networked to the Siemens 121C central monitor
Siemens | 1260 | Monitoring System | VGO-HXE | Fail | Fail | Pass | Pass | N/A | Pass | Pass | Yes | Yes | No | Unknown | Illawarra Area Health Service | 04/12/1997 | Although Siemens 121 central monitor reverts to 1900 after a few minutes the bedside monitrs are not affected, trending data was OK and recorder correctly printed the date
Siemens | Servo Ventilator 300 / Servo Screen 390 | Ventilator | PROM version 1.00 Servo screen V2.0 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No | N/A | Hunter Area Health Service | 06/05/1998 | Unit uses 4 digit year. Unit is Y2K compliant.
Siemens | Sicard 440 / Sicard 440S / Burdick 7500 | Stress ECG testing system |  | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | System uses 2 digit year. Unit is Y2K compliant (windowed compliance).
Siemens | Neurotron | Stimulator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Simed | S-100 | Monitor, Pulse Oximeter | 1.61 | Fail | Fail | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 08/10/1997 | Rollover 01/01/1900
Sony | PVM-2042QM | Monitor, Video | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions.
Sony | PVM-2043MD | Monitor, Video |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Sony | PVM-2053MD | Monitor, Video |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Sony | PVM-1440QM | Monitor, Video |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Spacelabs | 90303A | Monitor, Mainframe | 3.50.49 (1989) | Fail | [6 empty] | Yes | Yes | Yes | Contact Spacelabs for details | St Vincent's Hospital - Sydney | ??/??/1997 | S/W update avaliable. Minimum S/W version required is 3.67.04 (or 3.55.51 for Pre-Pauline models)
Spacelabs | 90385 (UCW) | Monitor, Mainframe | C1.02.33EN(BEHMOQ) | Pass | [6 empty] | No | N/A | Yes | N/A | St Vincent's Hospital - Sydney | ??/??/1997 | Unit is Y2K compliant.
Spacelabs | 90309 PC Scout | Monitor, Transport | 3.66.33 EN | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | N/A | Yes | N/A | Illawarra Area Health Service | 05/12/1997 | Unit is Y2K compliant.
Spacelabs | 4046 | Defibrillator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | St Vincent's Hospital - Sydney | 14/10/1997 | No programmable logic or real time clock. Unit is Y2K compliant.
Spacelabs | 90623A | Monitor, ECG, Press, Resp, Temp | 4.10.12 | Pass | Fail | Pass | Fail | N/A | Fail | Pass | Yes | Yes | Yes |  | Hunter Area Health Service | 11/02/1998 | Y2K non-compliance noted by Spacelabs (in Customer Bulletin). No update currently available. Contact Spacelabs for any further details.
Spacelabs | 90621A | Monitor, ECG, Resp | 4.10.20 | Fail | Fail | Pass | Pass | N/A | Pass | Pass | Yes | Yes | Yes |  | Hunter Area Health Service | 11/02/1998 | Y2K non-compliance noted by Spacelabs (in Customer Bulletin). No update currently available. Contact Spacelabs for any further details.
Spacelabs | 90604A | Monitor, ECG, Temp | 4.10.36 | Fail | Fail | Pass | Pass | N/A | Pass | Pass | Yes | Yes | Yes |  | Hunter Area Health Service | 11/02/1998 | Y2K non-compliance noted by Spacelabs (in Customer Bulletin). No update currently available. Contact Spacelabs for any further details.
Spacelabs | 90470 | Monitor, ECG, Press, Temp | V1.05.03 | Fail | Not tested | Not tested | Not tested | Not tested | Not tested | Not tested | Yes | Yes | Yes | Contact Spacelabs for details | Hunter Area Health Service | 11/02/1998 | S/W update avaliable. Minimum S/W version required is 2.02.09
Spacelabs | 90305 (PC2) | Monitor, Mainframe | V3.66.92EN | Fail | Fail | N/A | N/A | N/A | N/A | Pass | Yes | Yes | Yes | Contact Spacelabs for details | Hunter Area Health Service | 11/02/1998 | S/W update is currently under development. Contact Spacelabs for any further details.
Spacelabs | 90601A | Monitor, ECG | 4.10.36 | Fail | Fail | Pass | Pass | N/A | Pass | Pass | Yes | Yes | Yes |  | Hunter Area Health Service | 17/02/1998 | Y2K non-compliance noted by Spacelabs (in Customer Bulletin). No update currently available. Contact Spacelabs for any further details.
Spacelabs | 90426 | Module, NIBP, Adult / Infant | 1.02.34 | Pass | [9 empty] | N/A | Hunter Area Health Service | 03/04/1998 | Tested unit with compliant monitor (90309). No evidence of non-compliance. Unit is Y2K compliant.
Spacelabs | 90303B | Monitor, Mainframe | 3.51.46EN | Fail | Fail | Fail | Fail | N/A | Fail | Not tested | Yes | Yes | Yes | Contact Spacelabs | Hunter Area Health Service | 24/03/1998 | S/W update avaliable. Minimum S/W version required is 3.67.04 (or 3.55.51 for Pre-Pauline models)
Spacelabs | 90449 | Module, Printer |  | Pass |  | Pass |  | N/A | Pass | Pass | No | N/A | Yes | N/A | Hunter Area Health Service | 25/03/1998 | Unit is Y2K compliant.
Spacelabs | 90404 | Module, Cardiac Output | 1.11.02 | Pass | [6 empty] | Yes | Yes | Yes |  | Hunter Area Health Service | 25/03/1998 | Spacelabs claims non-compliance. Tested unit with compliant monitor (90309) and could not determine any non-compliance. Manual claims that date & time of each recording is stored though it only displays DAY & time. Tested OK.
Spacelabs | 90309 PC Scout | Monitor, Transport | 3.68.07EN | Pass | Pass | Pass | Pass | N/A | Pass | Pass | Yes | Yes | Yes | N/A | Hunter Area Health Service | 25/03/1998 | Unit is Y2K compliant.
Spacelabs | 90430 | Module, NIBP | 1.01.21 | Pass | [6 empty] | Yes | Yes | Yes | N/A | Hunter Area Health Service | 25/03/1998 | Tested unit with compliant monitor (90309). No evidence of non-compliance. Unit is Y2K compliant.
Spacelabs | 90419 | Module, Pulse Oximeter |  | Pass | [6 empty] | Yes | Yes | Yes | N/A | Hunter Area Health Service | 25/03/1998 | Tested unit with compliant monitor (90309). No evidence of non-compliance. Unit is Y2K compliant.
Spacelabs | 90603A | Monitor, Multiparameter | 4.10.20 | Fail | Fail | [5 empty] | Yes | Yes | Yes | Contact Spacelabs | Hunter Area Health Service | 27/03/1998 | Y2K non-compliance noted by Spacelabs (in Customer Bulletin). No update currently available. Contact Spacelabs for any further details.
Spacelabs | 90385 (UCW) | Monitor, Multiparameter | 1.04.33 | Pass | Pass | Pass | Pass | N/A | Pass | Pass | No | Yes | Yes | N/A | Hunter Area Health Service | 08/05/1998 | Unit uses 4 digit year. Unit is Y2K compliant (full compliance).
Spacelabs | 90467 | Module, NIBP, SaO2 | 1.04.02EN | Pass | Pass | Pass | Pass | N/A | Not tested | Not tested | No | Yes | Yes | N/A | Hunter Area Health Service | 08/05/1998 | Tested units with compliant moniitor (UCW). Checked tabular trends before and after rollover dates. Unit is Y2K compliant (full compliance).
Spacelabs | 90402 | Module, Pressure | 1.05.16 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date functions utilised by module. Unit is Y2K compliant.
Spacelabs | 90308 Express | Monitor, Transport | 3.51.05 | Fail | Fail | Fail | Fail |  |  |  | Yes | Yes | Yes | Contact Spacelabs | Hunter Area Health Service | 27/03/1998 | S/W update is currently under development. Contact Spacelabs for any further details.
Spacelabs | 90518 | Analyser, Multigas | 1.01.01EN | Pass | Pass | Pass | Pass |  |  |  | N/A | N/A | Yes | N/A | Hunter Area Health Service | 22/04/1998 | Tested unit with compliant monitor (90309). Unit is Y2K compliant.
Spacelabs | 90514 | Analyser, Multigas |  | Pass | Pass | Pass | Pass |  |  |  | N/A | N/A | Yes | N/A | Hunter Area Health Service | 22/04/1998 | Tested unit with compliant monitor (90309). Unit is Y2K compliant.
Spacelabs | 90513 | Analyser, CO2 |  | Pass | Pass | Pass | Pass |  |  |  | N/A | N/A | Yes | N/A | Hunter Area Health Service | 22/04/1998 | Tested unit with compliant monitor (90309). Unit is Y2K compliant.
Springfield Electronics | IIA | Monitor, Respiration, Apnoea | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Stackhouse | ST-1275 | Smoke Filtration System | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Storz | 482D | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Storz | 600DA | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes |  | Hunter Area Health Service | 23/02/1998 | No date / time function. No evidence of date exchange through TTL connector.
Storz | 485 | Light Source | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 13/03/1998 | No date / time functions.
Storz | 450V | Light Source |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 07/04/1998 | No date functions. Unit is Y2K compliant.
Stryker | 203 | Pump, Pressure, Irrigation | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Stryker | SE5 | Arthroscopic Surgical Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Stryker | OrthoLav | Pump, Pressure, Irrigation |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Taema | FC-10 | Spirometer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Telectronics | DS610 | Defibrillator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 06/03/1998 | No date functions.
Telectronics | DS610S | Defibrillator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 06/03/1998 | No date functions.
Telectronics | TE200 | Pacemaker, External | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/03/1998 | No date functions.
Telectronics | 174 | Pacemaker, Programmer |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Telectronics | 17 Series | Pacemaker, Programmer, Charger |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Teledyne | TED 200T | Monitor, Oxygen | Unknown | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | No | N/A | Hunter Area Health Service | 12/11/1997 | No date / time functions.
Teledyne | TED 100 | Monitor, Oxygen | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Terumo | UTD-7 | Monitor, Ultrasound, Fetal | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
TGS | US5 | Ultrasonic Scanner | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 27/03/1998 | No date functions.
Timeter | MP500 | Nebuliser, Ultrasonic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Toitu | N | Pump, Suction | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | No date / time functions.
Toronto Medical | L2 | Limb Manipulator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Tuta | 360-200 | Blood Warmer |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 12/11/1997 | No programmable logic or real time clock.
Tuta | Tuta | Water Bath | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date / time functions.
Ulco | VA1-10 | Ventilator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | No | N/A | Hunter Area Health Service | 13/11/1997 | No date / time functions.
Ulco | Elite 615 | Anaesthetic Machine | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | No |  | Hunter Area Health Service | 03/10/1997
Ulco | Event | Ventilator, Anaesthetic |  | Pass | Pass | N/A | N/A | N/A | Pass | N/A | Yes | Yes | No |  | St Vincent's Hospital - Sydney | 24/11/1997 | Real time clock data supported by CPU as 2 digits. Rollover 01/01/00. Date is only used on event recorder printout. (advice from manufacturer)
Ulco | EV-500 | Ventilator, Anaesthetic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Yes | Yes | No |  | St Vincent's Hospital - Sydney | 20/10/1997 | No real time clock. (advice from manufacturer)
Ulco | MK5, VA1-10 | Ventilator, Anaesthetic (with monitor / alarm) | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 11/02/1998 | No date functions.
Ulco | 5B MRI | Ventilator, Anaesthetic | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 23/02/1998 | Gas Unit - No date / time functions.
Ulco | AUS-PAP | CPAP Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | Gas Unit - No date / time functions.
Ulco | Compact | Anaesthetic Machine | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 23/04/1998 | No date functions. Unit is Y2K compliant.
Valley Forge Scientific | Codman CMC II | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Valleylab | SSE-2K | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 09/10/1997 | No computer technology or real time clock.
Valleylab | SSE-2L | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | St Vincent's Hospital - Sydney | 09/10/1997 | No computer technology or real time clock.
Valleylab | Force 2 | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 04/03/1998 | No date functions.
Valleylab | Surgistat | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 06/03/1998 | No date / time functions.
Valleylab | Force 40 | Electrosurgical Unit |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Valleylab | SSE3 | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Vickers Medical | 183 | Warmer, Overhead, Infant | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | No date / time functions.
Vitalograph | R | Spirometer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 12/03/1998 | No date / time functions.
Wallach | WA2000 | Monitor, Temperature, Opthalmic Probe | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions.
Wallach | WA3000 | Opthalmic Probe | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Walter Graphtek | PLEEG | Paperless EEG System | B 2.23 | Fail | Fail | [5 empty] | Yes | No | Yes |  | Hunter Area Health Service | ??/11/1996 | Date goes to 1980 at midnight on 31 Dec 1999 - unable to retrieve any patient information for patients with DOB >1979
Weck | 174701 | Insufflator | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date functions.
Welch Allyn | 74718 | Diagnostic Set | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Welch Allyn | 79156 | Head Set, Operating Lamp | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Welch Allyn | 61000 | Spirometer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Welch Allyn | 76106 | Printer / Charger, Spirometer | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 18/03/1998 | No date / time functions.
Wisap | 7110 | Insufflator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Wolf | 2046 | Insufflator |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No | N/A | Hunter Area Health Service | 08/05/1998 | No date functions. Unit is Y2K compliant (full compliance).
Zencor | MF1 | Electrosurgical Unit | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 05/03/1998 | No date / time functions.
Zimmer | ATS 500 | Pump, Air, Tourniquet | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No |  | Hunter Area Health Service | 18/03/1998 | No date / time functions.